LENSOMETER V2

R&D Tracker, Breakthrough Log & Technical Decision Record

Living project reference

| Item | Current value / note |
| --- | --- |
| Active repository | moeotech/lensometer-v2 |
| Primary benchmark | EyeNetra Lensometer / Netrometer methodology |
| Current experimental branch | V4 + V5 + V6 structured deflectometry |
| Known physical test lens | SPH -0.75 D, CYL -1.00 D |
| Document purpose | Track every code change, physical test, failure, breakthrough, equation, and next action so the project never loses direction. |
| Document date | 2026-09-01 |

Rule: every important test must enter this document with the exact setup, code commit, raw output, interpretation, and next decision.

# 1. Project Goal

Build a smartphone-based lensometer capable of estimating spectacle-lens refractive power from controlled image deformation of a known target, ultimately outputting SPH, CYL and AXIS. The project is inspired by deflectometry methods such as EyeNetra's Netrometer, but is being developed with our own target, geometry, image processing and calibration.

# 2. Core Measurement Concept

The core experiment is a two-image optical comparison: first capture the known target without a lens, then capture the same target through the lens. After registration, the software measures local displacement / spacing change, estimates directional deformation, fits principal components, and later converts the optical signal to diopters using device-specific calibration.

Reference target → lens target → correspondence → local deformation → principal directions → calibration → SPH / CYL / AXIS

# 3. What EyeNetra Taught Us

A true zero/reference image is fundamental.

A known target and fixed geometry simplify correspondence enormously.

The useful quantity is not merely image magnification; it is structured deflection / local deformation.

Calibration against known lenses is legitimate and expected.

Mechanical repeatability and controlled geometry matter as much as the software.

The final optical quantities should come from directional behavior: mean/isotropic power, anisotropy, and principal orientation.

# 4. Current System Status

| Component | Current state | Evidence | Action |
| --- | --- | --- | --- |
| Target detection | WORKING | Hundreds of points detected with revised dense dot target | Keep current target while testing |
| V5 correspondence | WORKING WELL | Up to 741 accepted matches and 97.6% spatial coverage | Freeze unless regression appears |
| V4 field estimation | WORKING / UNCALIBRATED | Stable raw λ1 / λ2 signals in successful runs | Use as comparison channel |
| V6 structured deflectometry | PARTIAL | Can succeed with fallback ROI; fails when ROI/cell coordinate handling breaks | Fix coordinate-space / ROI membership |
| SPH / CYL / AXIS output | NOT YET CALIBRATED | Experimental power remains N/A | Begin calibration after V6 is stable |
| Hardware geometry | EXPERIMENTAL | Phone stand + target + movable lens distance | Standardize distances before calibration |

# 5. Pipeline Evolution

## V4 — Direct Lens Optical Field

Uses geometric registration/correspondence to estimate a deformation field and principal values λ1 and λ2. It demonstrated that the lens signal can be extracted, but earlier tests were starved by weak topology/correspondence.

## V5 — Hardened Geometric Correspondence

Reworked matching independently of the old fragile grid topology. Uses robust registration, seed matches, RANSAC/inliers, expansion, residual filtering and spatial coverage. This became the first consistently strong matching layer.

## V6 — Structured Deflectometry

Attempts to mimic the physically meaningful route more closely: outside-lens registration, inside-lens local directional vectors, robust filtering, sinusoidal fitting and principal orientation. Its present blocker is ROI / coordinate-space reliability, not lack of target information.

# 6. Breakthrough Timeline

| Stage | Breakthrough / issue | What it taught us |
| --- | --- | --- |
| Early tests | Sparse / previous paper target | Topology collapsed to only a handful of usable correspondences. Optical math was being fed almost no spatial data. |
| EyeNetra study | Methodology breakthrough | Confirmed two-image deflectometry, known pattern, zero reference, best-fit deformation and calibration curve are the right family of methods. |
| V5 introduction | Software breakthrough | Topology-independent correspondence dramatically improved matching and spatial coverage. |
| Dense target redesign | Physical-target breakthrough | Changing the paper target produced hundreds of detected / matched points. V5 reached ~462 matches at 96.7% coverage in one test. |
| V6 first successful structured fit | Structured-deflectometry breakthrough | Example: 369 ref cells, 360 lens cells, 285 directional vectors, directional consistency PASS. |
| ROI failure discovery | Bug isolation | AUTO ROI produced impossible centerY=-172 and huge radius, proving ROI sanity validation was required. |
| Fallback ROI success | Recovery | Fallback ROI at (320,240), Rin=120, Rout=192 produced 35 common cells and 136 accepted vectors with PASS. |
| Latest decisive test | Coordinate-space diagnosis | V5 reached 741 accepted matches / 97.6% coverage while V6 reported zero cells inside fallback ROI. This strongly isolates a coordinate-system / ROI membership mismatch. |

# 7. Important Test Snapshots

## 7.1 Dense target, strong V5 / V6 run

| Item | Current value / note |
| --- | --- |
| V5 | 478 reference points, 473 lens points, 462 accepted inliers |
| V5 residual | Median 1.28 px |
| V5 coverage | 96.7% |
| V6 | 369 assigned ref cells, 360 assigned lens cells, 285 valid directional vectors |
| V6 A0 | 1.0068 |
| V6 amplitude | 0.0158 |
| V6 raw principal orientation | 116.9° |
| Interpretation | Pattern and correspondence were no longer the main blocker. |

## 7.2 Lens moved farther from camera

| Item | Current value / note |
| --- | --- |
| Known lens | SPH -0.75 D / CYL -1.00 D |
| V5 | 233 accepted inliers, 50.8% coverage, RMS 3.62 px |
| V6 | 186 valid directional vectors, directional consistency PASS |
| V6 A0 | 1.0024 |
| V6 amplitude | 0.0207 |
| V6 principal orientation | 12.7° |
| Interpretation | Lens-to-camera / lens-to-target geometry clearly changes measured optical signal and must be fixed before calibration. |

## 7.3 Fallback ROI successful run

| Item | Current value / note |
| --- | --- |
| ROI source | FALLBACK |
| ROI | Center (320,240), inner radius 120, outer radius 192 |
| Inside common cells | 35 |
| Accepted vectors | 136 / 245 |
| Accepted ratio median | 0.9989 |
| Ratio MAD | 0.0058 |
| A0 | 0.9991 |
| Amplitude | 0.0030 |
| Raw principal orientation | 78.3° |
| Directional consistency | PASS |

## 7.4 Latest decisive coordinate-space failure

| Item | Current value / note |
| --- | --- |
| V5 accepted matches | 741 |
| V5 spatial coverage | 97.6% |
| V5 transform | Tx 0.6, Ty -6.4, Rot -0.1°, Scale 0.995 |
| V5 RMS | 1.58 px |
| V6 ROI | Fallback center (320,240), Rin 120, Rout 192 |
| V6 inside ref / lens / common | 0 / 0 / 0 |
| V6 accepted vectors | 0 |
| Diagnosis | V6 cells and ROI are likely being compared in different coordinate systems. This is now the highest-priority software bug. |

# 8. Current Highest-Priority Blocker

V6 can possess hundreds of assigned/common grid cells in its debug visualization while the ROI membership stage reports zero cells inside a physically central fallback ROI. Because V5 succeeds on the same frames, this is not primarily a camera, target or lens-detection failure.

Current root-cause hypothesis: ROI center/radius and V6 cell coordinates are not expressed in the same canonical image coordinate system.

## Required verification telemetry

Analysis image width / height

Raw reference min/max X/Y

Assigned reference min/max X/Y

Assigned lens min/max X/Y

Common-cell min/max X/Y

ROI center X/Y and radii

Nearest assigned/common-cell distance to ROI center

Exact coordinate transforms used for crop, rotation, scale and preview rendering

# 9. Optical Quantities We Are Tracking

| Quantity | Meaning in our pipeline | Status |
| --- | --- | --- |
| λ1 / λ2 | Principal directional deformation values from fitted local field | Raw signal available |
| Isotropic component | Average / common deformation component, conceptually related to spherical power | Raw signal available |
| Anisotropic component | Difference between principal directions, conceptually related to cylinder | Raw signal available |
| Principal angle | Orientation of the directional deformation | Raw signal available |
| A0 | Mean directional ratio in V6 sinusoidal fit | Available |
| Amplitude | Magnitude of directional anisotropy in V6 | Available |
| SPH | Calibrated sphere result in diopters | Not yet calibrated |
| CYL | Calibrated cylinder result in diopters | Not yet calibrated |
| AXIS | Calibrated cylinder axis | Not yet validated |

# 10. Calibration Plan

Do not attempt final diopter conversion until the physical geometry and V6 ROI/matching are repeatable. Once stable, measure a controlled set of known lenses and fit the raw optical signals to known SPH/CYL values.

| Lens type | Known power | Purpose |
| --- | --- | --- |
| Plano | 0.00 D | Baseline |
| Sphere | -0.50 D | Negative low power |
| Sphere | -1.00 D | Negative medium |
| Sphere | -2.00 D | Negative stronger |
| Sphere | +0.50 D | Positive low |
| Sphere | +1.00 D | Positive medium |
| Cylinder | -0.50 D at known axis | Anisotropy / axis calibration |
| Cylinder | -1.00 D at known axis | Anisotropy / axis calibration |
| Current known lens | SPH -0.75 D, CYL -1.00 D | Primary development check lens |

# 11. Standard Physical Test Protocol

Use the same printed target revision and exact print scale.

Fix the phone position and camera orientation.

Record camera-to-lens distance.

Record lens-to-target distance.

Capture reference without moving the phone or target.

Insert the lens into the defined plane and center it consistently.

Run V5 and V6 on the same frames.

Export both raw logs.

Record the Git commit SHA before the test.

Repeat each condition at least 3 times before interpreting optical trends.

# 12. Experiment Log Template

Duplicate this section for every important physical test.

| Item | Current value / note |
| --- | --- |
| Experiment ID | EXP-YYYYMMDD-### |
| Date / time |  |
| Git commit SHA |  |
| Phone / device |  |
| Target version |  |
| Known lens |  |
| Lens orientation / axis marking |  |
| Camera → lens distance |  |
| Lens → target distance |  |
| Lighting |  |
| ROI source |  |
| V5 accepted matches |  |
| V5 coverage |  |
| V5 RMS |  |
| V6 assigned ref / lens / common |  |
| V6 accepted vectors |  |
| V6 ratio median / MAD |  |
| V6 A0 |  |
| V6 amplitude |  |
| V6 raw principal orientation |  |
| V4 λ1 / λ2 |  |
| Observed SPH / CYL / AXIS output |  |
| Result | PASS / PARTIAL / FAIL |
| Interpretation |  |
| Decision |  |
| Next action |  |

# 13. Code Change Log Template

| Date | Commit | Area | Change | Why | Result |
| --- | --- | --- | --- | --- | --- |

# 14. Decision Log

| Decision | Reason | Status |
| --- | --- | --- |
| Keep V4 | Useful independent raw-field comparison channel. | ACTIVE |
| Keep V5 | Current strongest correspondence engine. Do not rewrite while it is producing high coverage. | ACTIVE |
| Continue V6 | Closer to structured deflectometry / EyeNetra-style measurement logic. | ACTIVE |
| Do not calibrate yet | Geometry and ROI pipeline must first be stable. | ACTIVE |
| Use dense dot target | Massively improved feature count and correspondence. | ACTIVE |
| Treat hardware as part of the instrument | Distances and lens plane affect the measured signal and must be repeatable. | ACTIVE |
| Do not invent EyeNetra internals | Use only what their public code/papers actually establish. | ACTIVE |

# 15. Research References

EyeNetra Lensometer Android repository: https://github.com/vitorpamplona/EyeNetraLensometer

EyeNetra Netrometer paper: https://eyenetra.com/studies/Frontiers%20In%20Optics%20-%20Netrometer.pdf

EyeNetra Netrometer product page: https://eyenetra.com/product-netrometer.html

The Framery Lens Scanner: https://www.theframery.com/lens-scanner

Our active repository: moeotech/lensometer-v2

# 16. Immediate Next Actions

| Priority | Action | Done |
| --- | --- | --- |
| P0 | Fix V6 coordinate-space / ROI membership bug. | ☐ |
| P0 | Confirm V6 identifies >20 realistic common cells in the central lens ROI on the strong 97.6%-coverage test. | ☐ |
| P1 | Freeze physical geometry for repeatability. | ☐ |
| P1 | Repeat the known -0.75 SPH / -1.00 CYL lens at fixed geometry, 5-10 runs. | ☐ |
| P1 | Check repeatability of A0, amplitude, principal angle, λ1 and λ2. | ☐ |
| P2 | Create first calibration table using known spherical lenses. | ☐ |
| P2 | Add known cylindrical lenses at marked axes to validate axis behavior. | ☐ |
| P3 | Only after calibration, enable experimental SPH/CYL/AXIS output. | ☐ |

This document should be updated after each meaningful Git commit or physical test. Keep raw exports attached separately when they are too large for the tracker.